## Slide 1: Quizalofop-P-ethyl CIPAC Report （FULL SCALE COLLABORATIVE STUDY）
Anhui Fengle Agrochemical Co., Ltd.
May 2020
Fengle Agrochem

## Slide 2: CONTENT
1. PARTICIPANTS
2. GENERAL INFORMATION
3. SAMPLES
4. METHOD
5. REMARKS OF THE PARTICIPANTS
6. EVALUATION AND DISCUSSION
7. CONCLUSIONS
Fengle Agrochem

## Slide 3: PARTICIPANTS
Fengle Agrochem
Name of responsible person | Lab Name | City, Country
Quanzhi Deng | Shanghai pesticide research institute Co., Ltd. | Shanghai, China
Peng Wu | National Center for Pesticide Quality Supervision and Inspection (Beijing), Institute for the Control of Agrochemicals, Ministry of Agriculture and Rural Affairs , P. R. China | Beijing, China
Liang Wu | ZHEJIANG TIANFENG BIOSCIENCE CO., LTD. | Jinhua, China
Liang Huang | Anhui Fengle Agrochemical Co., Ltd. Product Testing Center | Hefei, China
Xiu’e Yin | Pesticides Test Laboratory Shenyang Research Institute of Chemical Industry Co., Ltd | Shenyang, China
Angela Santilio | Italian National Institute of Health - Istituto Superiore di Sanità (ISS) | Rome, Italy

## Slide 4: PARTICIPANTS
Fengle Agrochem
Name of responsible person | Lab Name | City, Country
Ivan Orgei | FRANDESA Co. LTD | Minsk, Belarus
Xiaoying Liu | Laprode (ZheJiang) Analysis Co., Ltd. | Hangzhou, China
Hiroka Harada | Agricultural Chemicals Inspection Station Food and Agricultural Materials Inspection Center | Tokyo, Japan
Juntao Zhang | Helm China Co., Ltd. | Shanghai, China
Wagener Peter | Bayer CropScience Division | Frankfurt-Höchst, Germany
Aysel Takkabulan | Ankara Plant Protection Central Research Institute Pesticide Control Laboratuary Ankara Directorate of Plant Protection Central Research Institute | Ankara, Turkey

## Slide 5: PARTICIPANTS
Fengle Agrochem
Name of responsible person | Lab Name | City, Country
Maria Chiara Neri | ChemService S.r.l. Controlli e Ricerche | Novate Milanese (MI)Italy
Robin Xiao | Jiangsu BioGuide Laboratory Co., Ltd. | Changzhou, China
Madhuri | BIONEEDS INDIA PRIVATE LIMITED | Karnataka, India
Agus Salim | Laboratorium PT Agricon | West Java, Indonesia
Yusuf Vohra | Jai Research Foundation | Gujarat, India
Participants are listed in random sequence, lab numbers in the result tables were assigned in sequence of result receipt.

## Slide 6: GENERAL INFORMATION
Chemical name: Ethyl (R)-2-[4-(6-Chloroquinoxalin-2-yloxy)phenoxy]propionate
ISO common name: Quizalofop-P-ethyl
CAS No.:100646-51-3
Structure:
Molecular mass: 372.8
M.p. 76.1~77.1 °C
Empirical formula: C19H17ClN2O4
Solubility:
In water, 6.1×10-4 g/l at 20 °C, pH 5.0 – 7.0; Xylene, ethyl acetate and acetone > 250 g/l,
1,2-dichloroethane > 1000 g/l at 22 – 23 °C; methanol 34.87 g/l, n-heptane 7.168 g/l at 20 °C
Description: Off-white powder
Stability: Stable at neutral and acidity condition.
Fengle Agrochem

## Slide 7: SAMPLES
The participants who completed the study are listed in section 1.
Five test samples, the analytical standard were sent to the participants:
1. TC (Sample A)
2. TC (Sample B)
3. EC formulation (Sample C)
4. EC formulation (Sample D)
5. EC formulation (Sample E)
Quizalofop-p-ethyl analytical standard, 99.3 % purity.
17 participants sent back their results in time.
Fengle Agrochem

## Slide 8: METHOD
1. Scope
The content of Quizalofop-p-ethyl is determined in technical material and emulsifiable concentrate products.
2.Principle
The Quizalofop-p-ethyl content of the samples is determined by normal phase HPLC on chiral column using UV detector at 237nm and external standardization.
3.Procedure
The samples were analyzed on two different days, each day involving duplicate injections of duplicate weights. Both test and reference solutions were freshly prepared on each day.
Fengle Agrochem

## Slide 9: REMARKS OF THE PARTICIPANTS
Several participants made comments about the performance of the method and noted deviations from the method:
Fengle Agrochem
Laboratory | Column | Remarks
Lab 1 | Diacel, Dimensions: 4.6 x 250mm, Chiralcel AD-H, 5µm, Serial No: ADH0CE-XI076 | Injection volume changed from 1.5µL to 3µL.
Lab 2 | Chiralcel AD-H, 250mm x 4.6 mm (id), 5μm, serial No.: ADOHCE-XH017 | None
Lab 3 | Diacel, Chiralcel OD-H, 250mm x 4.6 mm (id), 5μm, serial No.: ODH0CE-LJ097 | Column: Chiralcel AD-H was replaced by Chiralcel OD-H. Mobile phase: n-heptane+isopropanol =99:1 was used in the revised method. Injection volume: 5μL was injected in the revised method.

## Slide 10: REMARKS OF THE PARTICIPANTS
Fengle Agrochem
Laboratory | Column | Remarks
Lab 4 | CHIRALPAK AD-H, DAICEL, 4.6mm*250mm, Lot No.ADH0CE-QD122 | None
Lab 5 | OD-H chiralcel DAICEL/4.6mm/ ODH0ce-SB018 | Column: OD-H chiralcel, Mobile phase: n-heptane: isopropanol =94:6(v/v), Retention time: approximately 15.6 min.
Lab 6* | DAICEL ODH 250mm * 0.46 | None
Lab 7 | CHIRALPAK® AY-H 250 x 4.6 mm x 5 μm--- Amylose tris(5-chloro-2-methylphenylcarbamate) | None
Lab 8 | DAICEL, CHIRALCEL OD-H, Dimensions: 4.6 x 250 mm, Particle Size: 5μm, Serial No. : ODH0CE-RJ034 | None

## Slide 11: REMARKS OF THE PARTICIPANTS
Fengle Agrochem
Laboratory | Column | Remarks
Lab 9 | CHIRALPAK AD-H (Daicel, 4.6 mm x 250 mm, 5 µm) | Injection volume can’t be set 1.5 µL due to apparatus specifications. Column temperature 30°C.
Lab 10 | CHIRALPAK AD-H 250mm x 4.6mm, 5μm, Part No. 19325 | Injection volume 1.5 μL is changed to 5 μL
Lab 11 | Chiralcel AD-H 5µm, Daicel, 250 mm x 4.6mm, Part No. 19325 | increased the flow from 0.6 to 0.8 mL/min
Lab 12 | Dimensions: 250mm x 4.6 mm, CHIRALPAK AD-H, Part No: 19325 Particle Size: 5 μm | Run-time for each injection was determined as 20 minutes.
Lab 13 | Daicel Corporation, Chiralpak AD-H, 4.6 mm x 250 mm, 5 µm. Lot number ADH0CE-XH094 | None
Lab 14 | CHIRALPAK, AD-H, 250 x 4.6mm x 5um, Part No.: 20325 | None

## Slide 12: REMARKS OF THE PARTICIPANTS
Fengle Agrochem
Laboratory | Column | Remarks
Lab 15 | Daicel, Dimensions: 250 x 4.6mm, 5 µm, Serial Number:19325 | Injection volume changed to 5 µL Instead of 1.5 µL
Lab 16 | Phenomenex Lux® 5 μm Amylose-2, Dimension: 250 x 4.6 mm (i.d), S/No: H16-374037, Part No: OOG-4472-EO | Column HPLC used not similar column in reference method. Isopropanol change with ethanol, to get better peak chromatogram; diluting the solution, both standard and sample, 10 times, and the volume of the injection became 50µL. Flow rate was changed from 0.6 to 0.8 ml/min.
Lab 17 | Chiralpak AD-H, 250 mm x 4.6 mm (i.d), 5.0 µ particle size | Injection volume was changed from 1.5 µL to 2.0 µL.
*: The data of Lab 6 is not usable. Because the chromatographic column was changed to DAICEL ODH 250mm * 0.46, the mobile phase was not adjusted accordingly, the enantiomers were not separated completely.

## Slide 13: EVALUATION AND DISCUSSION
  Evaluation of the Quality of Data and Chromatograms
The data obtained from each of the laboratories were reviewed to determine if there were any significant deviations regarding the chromatography which might affect the analysis results.
Fengle Agrochem

## Slide 14: EVALUATION AND DISCUSSION
  Determination of Quizalofop-p-ethyl
Results reported by the laboratories and the statistical evaluation are listed in tables 1-5 and displayed in figures 1-5.
The statistical evaluation of the data was done following the “Guidelines for CIPAC Collaborative Study Procedures for Assessment of Performance of Analytical Methods”, according to DIN ISO 5725. The testing for outliers/stragglers of the laboratory means values were performed according to Grubbs test on a 1%/5% significance level, respectively. The Grubbs test identified stragglers and outliers for the EC formulations as well as for the technical material (marked with +/++ in Table 2).
Fengle Agrochem

## Slide 15: EVALUATION AND DISCUSSION
  Determination of Quizalofop-p-ethyl
All results reported by the 16 laboratories are reported and the statistical evaluation of these are listed in Tables 1-3 and displayed in Figures 1-5(The data of Lab 9 is not usable). These results are reported without any exclusion of outliers and/or stragglers.
In addition, a separate evaluation, listed in Table 4-5, display the results with the exclusion of stragglers and outliers.
A comparison of the RSDR of this collaborative study with the unmodified Horwitz equation showed that the relative reproducibility standard deviation (RSDR) is below the Horwitz value in all five samples even without elimination of stragglers and outliers (see Table 3). The RSDR is further improved if stragglers and outliers are removed (see Table 4 and 5). No more than one value has been removed per sample (Table 5). The validity of the results and the suitability of the analytical method are shown. This collaborative trial is acceptable.
Fengle Agrochem

## Slide 16: EVALUATION AND DISCUSSION
Fengle Agrochem
 | Sample A |  | Sample B |  | Sample C |  | Sample D |  | Sample E | 
 | Day1 | Day2 | Day1 | Day2 | Day1 | Day2 | Day1 | Day2 | Day1 | Day2
Lab 1 | 993.1 | 994.3 | 992.0 | 988.6 | 102.8 | 102.2 | 105.1 | 104.4 | 104.3 | 104.8
Lab 2 | 983.4 | 984.5 | 987.3 | 986.5 | 98.4 | 98.7 | 99.6 | 100.3 | 98.6 | 98.5
Lab 3 | 994.4 | 992.2 | 988.2 | 988.8 | 99.6 | 98.8 | 99.5 | 99.9 | 98.5 | 98.5
Lab 4 | 990.0 | 990.1 | 991.3 | 990.3 | 99.0 | 98.8 | 100.3 | 100.1 | 99.6 | 99.6
Lab 5 | 990.9 | 989.8 | 988.8 | 988.6 | 101.4 | 101.1 | 102.2 | 102.2 | 102.1 | 101.8
Lab 6 | n.u.* | n.u.* | n.u.* | n.u.* | n.u.* | n.u.* | n.u.* | n.u.* | n.u.* | n.u.*
Lab 7 | 1004.8 | 985.8 | 980.5 | 978.9 | 100.2 | 101.3 | 105.4 | 103.0 | 103.7 | 102.2
Lab 8 | 995.6 | 993.6 | 997.9 | 992.1 | 100.5 | 99.8 | 101.5 | 101.5 | 102.8 | 101.8
Lab 9 | 1010.0 | 1011.0 | 1006.0 | 1001.0 | 99.0 | 100.0 | 101.0 | 101.0 | 100.0 | 100.0
Table1: Quizalofop-p-ethyl assay in TC and EC (g/kg); results for each laboratory
on day 1 and day 2

## Slide 17: EVALUATION AND DISCUSSION
Fengle Agrochem
 | Sample A |  | Sample B |  | Sample C |  | Sample D |  | Sample E | 
 | Day1 | Day2 | Day1 | Day2 | Day1 | Day2 | Day1 | Day2 | Day1 | Day2
Lab 10 | 983.3 | 987.7 | 992.6 | 993.2 | 97.0 | 98.5 | 99.0 | 99.8 | 96.9 | 98.4
Lab 11 | 992.4 | 992.3 | 995.1 | 997.8 | 99.0 | 99.3 | 100.5 | 100.9 | 99.7 | 100.0
Lab 12 | 978.9 | 985.4 | 1002.7 | 1002.7 | 99.1 | 99.6 | 100.7 | 101.3 | 99.3 | 101.1
Lab 13 | 986.2 | 981.6 | 987.1 | 983.4 | 99.0 | 98.4 | 99.9 | 99.1 | 100.5 | 99.7
Lab 14 | 980.6 | 982.7 | 992.0 | 991.0 | 98.7 | 98.9 | 102.1 | 101.1 | 100.3 | 100.5
Lab 15 | 986.9 | 987.8 | 991.2 | 990.9 | 100.3 | 100.1 | 100.1 | 100.1 | 101.5 | 102.0
Lab 16 | 987.8 | 993.8 | 990.7 | 988.7 | 99.1 | 100.5 | 100.2 | 101.3 | 100.3 | 102.0
Lab 17 | 988.2 | 985.4 | 986.9 | 986.4 | 101.4 | 102.8 | 102.7 | 103.2 | 102.3 | 103.3
* Result is not usable

## Slide 18: EVALUATION AND DISCUSSION
Fengle Agrochem
Table 2: Mean values
 | SAMPLE A | SAMPLE B | SAMPLE C | SAMPLE D | SAMPLE E
Lab 1 | 993.700 | 990.300 | 102.500 | 104.750 | 104.550
Lab 2 | 983.950 | 986.900 | 98.550 | 99.950 | 98.550
Lab 3 | 993.300 | 988.500 | 99.200 | 99.700 | 98.500
Lab 4 | 990.050 | 990.800 | 98.900 | 100.200 | 99.600
Lab 5 | 990.350 | 988.700 | 101.250 | 102.200 | 101.950
Lab 6 | n.u.* | n.u.* | n.u.* | n.u.* | n.u.*
Lab 7 | 995.300 | 979.700 | 100.750 | 104.200 | 102.950
Lab 8 | 994.600 | 995.000 | 100.150 | 101.500 | 102.300
Lab 9 | 1010.500+/++ | 1003.500 | 99.500 | 101.000 | 100.000

## Slide 19: EVALUATION AND DISCUSSION
Fengle Agrochem
 | SAMPLE A | SAMPLE B | SAMPLE C | SAMPLE D | SAMPLE E
Lab 10 | 985.500 | 992.900 | 97.750 | 99.400 | 97.650
Lab 11 | 992.350 | 996.450 | 99.150 | 100.700 | 99.850
Lab 12 | 982.150 | 1002.700 | 99.350 | 101.000 | 100.200
Lab 13 | 983.900 | 985.250 | 98.700 | 99.500 | 100.100
Lab 14 | 981.650 | 991.500 | 98.800 | 101.600 | 100.400
Lab 15 | 987.350 | 991.050 | 100.200 | 100.100 | 101.750
Lab 16 | 990.800 | 989.700 | 99.800 | 100.750 | 101.150
Lab 17 | 986.800 | 986.650 | 102.100 | 102.950 | 102.800
* Result is not usable
+ Grubbs Test straggler
++Grubbs Test outlier

## Slide 20: EVALUATION AND DISCUSSION
Fengle Agrochem
Table 3: Summary of the statistical evaluation with outliers /stragglers
 | sample A | sample B | sample C | sample D | sample E
Xm | 990.14 | 991.23 | 99.79 | 101.22 | 100.77
L | 16 | 16 | 16 | 16 | 16
Sr | 3.979 | 1.781 | 0.583 | 0.592 | 0.661
SR | 7.578 | 6.196 | 1.373 | 1.658 | 1.912
r | 11.141 | 4.987 | 1.632 | 1.658 | 1.851
R | 21.218 | 17.349 | 3.844 | 4.642 | 5.354
RSDr | 0.402 | 0.180 | 0.584 | 0.584 | 0.656
RSDR | 0.765 | 0.625 | 1.376 | 1.638 | 1.897
RSDR(Hor) | 2.003 | 2.003 | 2.829 | 2.823 | 2.825
HorRat | 0.382 | 0.312 | 0.486 | 0.580 | 0.671

## Slide 21: EVALUATION AND DISCUSSION
Fengle Agrochem
Table 4: Summary of the statistical evaluation without outliers
 | sample A | sample B | sample C | sample D | sample E
Xm | 990.45 | 991.23 | 99.79 | 101.22 | 100.77
L | 15 | 16 | 16 | 16 | 16
Sr | 4.030 | 1.781 | 0.583 | 0.592 | 0.661
SR | 7.716 | 6.196 | 1.373 | 1.658 | 1.912
r | 11.284 | 4.987 | 1.632 | 1.658 | 1.851
R | 21.605 | 17.349 | 3.844 | 4.642 | 5.354
RSDr | 0.407 | 0.180 | 0.584 | 0.584 | 0.656
RSDR | 0.779 | 0.625 | 1.376 | 1.638 | 1.897
RSDR(Hor) | 2.003 | 2.003 | 2.829 | 2.823 | 2.825
HorRat | 0.389 | 0.312 | 0.486 | 0.580 | 0.671
Sample A: Results of Lab 9 eliminated.

## Slide 22: EVALUATION AND DISCUSSION
Fengle Agrochem
Table 5: Summary of the statistical evaluation without stragglers
 | sample A | sample B | sample C | sample D | sample E
Xm | 990.45 | 991.23 | 99.79 | 101.22 | 100.77
L | 15 | 16 | 16 | 16 | 16
Sr | 4.030 | 1.781 | 0.583 | 0.592 | 0.661
SR | 7.716 | 6.196 | 1.373 | 1.658 | 1.912
r | 13.0256 | 27.2328 | 8.1928 | 4.8776 | 8.3748
R | 15.0864 | 27.4876 | 8.8116 | 9.4724 | 14.1988
RSDr | 0.407 | 0.180 | 0.584 | 0.584 | 0.656
RSDR | 0.779 | 0.625 | 1.376 | 1.638 | 1.897
RSDR(Hor) | 2.003 | 2.003 | 2.829 | 2.823 | 2.825
HorRat | 0.389 | 0.312 | 0.486 | 0.580 | 0.671
Sample A: Results of Lab 9 eliminated.

## Slide 23: EVALUATION AND DISCUSSION
Fengle Agrochem
Fig. 1:
Figures 1 – 5(all results except lab 6).
Chart: Quizalofop-p-ethy Sample A
| Lab 1 | Lab 2 | Lab 3 | Lab 4 | Lab 5 | Lab 6 | Lab 7 | Lab 8 | Lab 9 | Lab 10 | Lab 11 | Lab 12 | Lab 13 | Lab 14 | Lab 15 | Lab 16 | Lab 17
1 | 993.7175 | 983.9525 | 993.25 | 990.05 | 990.35 | 995.35 | 994.5625 | 1012.0226 | 985.475 | 992.3634 | 982.1475 | 983.875 | 981.635 | 987.3 | 990.8275 | 986.775
2 | 975.26 | 975.26
3 | 1005.19 | 1005.19
4 | 965.93 | 965.93
5 | 1014.53 | 1014.53
6 | 990.23 | 990.23

## Slide 24: EVALUATION AND DISCUSSION
Fengle Agrochem
Fig. 2:
Figures 1 – 5(all results except lab 6).
Chart: Quizalofop-p-ethy Sample B
| Lab 1 | Lab 2 | Lab 3 | Lab 4 | Lab 5 | Lab 6 | Lab 7 | Lab 8 | Lab 9 | Lab 10 | Lab 11 | Lab 12 | Lab 13 | Lab 14 | Lab 15 | Lab 16 | Lab 17
1 | 990.2875 | 986.9025 | 988.975 | 990.8 | 988.7 | 979.725 | 994.965 | 1003.5 | 992.875 | 996.4266 | 1002.6725 | 985.25 | 991.4875 | 991.05 | 989.7325 | 986.62
2 | 981.65 | 981.65
3 | 1000.85 | 1000.85
4 | 972.39 | 972.39
5 | 1010.11 | 1010.11
6 | 991.25 | 991.25

## Slide 25: EVALUATION AND DISCUSSION
Fengle Agrochem
Fig. 3:
Figures 1 – 5(all results except lab 6).
Chart: Quizalofop-p-ethy Sample C
| Lab 1 | Lab 2 | Lab 3 | Lab 4 | Lab 5 | Lab 6 | Lab 7 | Lab 8 | Lab 9 | Lab 10 | Lab 11 | Lab 12 | Lab 13 | Lab 14 | Lab 15 | Lab 16 | Lab 17
 | 102.5 | 98.555 | 99.15 | 98.875 | 101.225 | 100.75 | 100.1375 | 99.5 | 97.7 | 99.1436 | 99.3725 | 98.675 | 98.81 | 100.2 | 99.7825 | 102.0975
mean r | 97.6058 | 101.9534
 | 95.6431 | 103.9161
4 | 99.78 | 99.78

## Slide 26: EVALUATION AND DISCUSSION
Fengle Agrochem
Fig. 4:
Figures 1 – 5(all results except lab 6).
Chart: Quizalofop-p-ethy Sample D
| Lab 1 | Lab 2 | Lab 3 | Lab 4 | Lab 5 | Lab 6 | Lab 7 | Lab 8 | Lab 9 | Lab 10 | Lab 11 | Lab 12 | Lab 13 | Lab 14 | Lab 15 | Lab 16 | Lab 17
 | 104.79 | 99.985 | 99.7 | 100.175 | 102.225 | 104.25 | 101.515 | 101.5 | 99.35 | 100.7069 | 101.0325 | 99.525 | 101.57 | 100.1 | 100.75 | 102.9
mean r | 98.2459 | 104.2634
 | 96.0369 | 106.4724
4 | 101.25 | 101.25

## Slide 27: EVALUATION AND DISCUSSION
Fengle Agrochem
Fig. 5:
Figures 1 – 5(all results except lab 6).
Chart: Quizalofop-p-ethy Sample E
| Lab 1 | Lab 2 | Lab 3 | Lab 4 | Lab 5 | Lab 6 | Lab 7 | Lab 8 | Lab 9 | Lab 10 | Lab 11 | Lab 12 | Lab 13 | Lab 14 | Lab 15 | Lab 16 | Lab 17
 | 104.5275 | 98.535 | 98.475 | 99.575 | 101.95 | 102.975 | 102.3275 | 100.25 | 97.625 | 99.8516 | 100.2075 | 100.1 | 100.415 | 101.725 | 101.12 | 102.775
mean r | 98.493 | 103.0613
 | 95.2275 | 106.3268
4 | 101.25 | 101.25

## Slide 28: CONCLUSIONS
17 different laboratories participated in this collaborative study. The results of the labs are given in Table 1-2, the statistical summary given in Table 3-5. The results are illustrated in figures 1 – 5.
Without elimination of any outliers or stragglers the Relative Reproducibility Standard Deviation (% RSDR ) is below the calculated acceptable value based on the Horwitz curve calculation (% RSDR (Hor)) in all samples.
The number of considered results after elimination of stragglers and outliers was 15.
Taking into account the relatively high number of participating laboratories a broad basis was given even after elimination of the outliers. Therefore, we consider this method to be suitable without further changes and recommend accepting it as a provisional CIPAC MT-method for the determination of Quizalofop-p-ethyl in technical material and EC formulation.
Fengle Agrochem

## Slide 29: THANK YOU !
Fengle Agrochem